- Virtual ports software, click add ports to add virtual COM port, as following picture, add COM 1 & COM 2 as linked ports

- 125  software setting , here COM3 is for connect the 125 analyzer, COM1 is the virutal COM port for LIS simulator connection.

- The 125 should have following test items:

BUN, CREA, GLUC, CREA, URIC, CHOL, TG, HDL, LDL, TP, ALB, TB, DB, AST, ALT, ALP

The sample informaiton in the simulator:

Barcode: 112233,  test list: BUN, CREA, GLUC

Barcode: 445566,  Test List: GLUC, BUN, CREA, URIC, CHOL, TG, HDL, LDL, TP, ALB, TB, DB, AST, ALT, ALP

- the LIS interface simulator software: Open project file: PKL125-3.ptp

In the Project setting, select the COM 2 for communication, then click Start Communication

- in the 125 software:

SCHEDULE---> LIS

Click the Receive button, then input the PID:112233, click Save , then Receive

- in the simulation software:

Click the [Send] in following sequent

ACK  Header  P.Info 1  Order Combin 1  EOT  END 1

- in the 125 softare, will receive a PID:112233 test list, as following

Click Save to save this PID to the working list of 125.

The Steps above simulate download a PID:112233 test list from a server, and all the communication is record in the simulator software, that will help to develop a Lis interface.